# IV.	Post-Processing/Data Analysis

In order to provide science groups with preliminary data while still in the field, some real-time data analysis can be done on the vessel during the cruise.  However, this will only provide preliminary results as final results must await post-cruise sensor calibrations and configuration optimization.

The following commands can be run on any PC with the SEASOFT software installed and the data files located in that directory.  From the C:\SEASOFT> prompt, these commands will create .CNV files (converted data files), which will be converted to ASCII format for easier processing of the data.

| Command | Description |
| --- | --- |
| datcnv | Converts the data from SEASOFT raw binary data (.DAT). |
| alignctd | Aligns temperature, conductivity, and oxygen measurements in time relative to pressure. Eliminates any lags in data (from secondary sensors and oxygen). |
| wildedit | Eliminates any wild data points. The 911plus CTD unit should take care of this on its own, but running the command is a good backup. |
| celltm | Conductivity cell thermal mass correction. |
| filter | Simple filter to make data smoother (pressure resolution). |
| loopedit | Corrects when the CTD package is moving less than the minimum velocity or traveling backwards due to ship roll. |
| derive | Gives values for oxygen. The file will save as derive.cfg; rename this file to derox.cfg to avoid confusion with next derive configuration. |
| binavg | Averages data into the desired pressure bins (e.g., every meter). This reduces noise and smoothes data. Reduces the size of the file. |
| derive | Derives salinity, temperature, density, and other parameters. The file will save as derive.cfg. Rename this file to derts.cfg to avoid writing over other derive configurations. |
| rossum | Takes the averaged data at each bottle trip and prints pressure, bottle number, and other values in a .ROS file. This should be saved in a public folder for easy access by researchers who sample bottles. It should also be printed and placed in the log book. |
| asciiout | Converts all info from binary to ASCII format in an .ASC file. This should be kept in a public space as well. |

The LTER group has created a batch file in order to make processing this data easier.  See the computer technician on board to revise this batch file for your particular cruise.

- The batch file needs to be run from the C:\SEASOFT> directory. The batch file is located on the P:\ drive.  Copy the batch file to the C:\SEASOFT> directory at the beginning of the cruise.  At the C:\SEASOFT> prompt , type

postproc.bat  <3 digit cast number>  <scans to skip>

and press <Enter>.

- The batch file will run the SEASOFT commands as listed above, and output an ASCII file (.ASC) and a rosette file (.ROS) to P:\CTD\ASCII and P:\CTD\bottle, respectively.

- The LTER group also has a MATLAB script which will plot temperature and salinity for a particular station.  This script is also found in P:\CTD.

For more elementary manual processing, the only commands that need to be run are datcnv, binavg, derive, rossum, and asciiout.